## Master Plan
Design Requirements |  | Method of Verification (BOM, X type of testing, self-evident by dimension X, etc). | If you know there would be a method to test that you want to do, but in discussing with the instructional team know it may not be feasible within the class framework, make a secondary plan for a simulated testing procedure that may work towards the goal within the limits of the class in addition to the "If we could, we would do this" plan. |  | Reporting | Target Completion Date | Required Materials/Prep | Comments
|  |  | Ideal Testing Plan (Lay out plan for testing generally, as detailed as possible) | Preliminary Testing Plan (Lay out plan for testing generally, as detailed as possible)
1.1 | Safety/Sterilizability
| Bacterial Removal | Sterilization Testing | Add 1 mL of bacteria to each component and metal surgical tube. Incubate for 48 hours. Then wipe down with kimwipe for 15 seconds. Place component in autoclave. Once done, remove from autoclave, look under microscope for bacteria. | Add 1 mL of bacteria to each component and metal surgical tube, then wipe down with kimwipe for 15 seconds. Place component in autoclave. Once done, remove from autoclave, look under microscope for bacteria. | Determine and record % of surface area covered by bacteria at 5 different locations for both surgical-grade metal and each component | 2024-11-08 | Equipment and instruments needed: Autoclave, Each component of our device, Surgical-grade metal, Light microscope, Bacteria stock of Escherichia coli Reusable glassware/plasticware needed: Eppendorf tubes (1.5mL) x12 Disposable materials needed: Pipettes (1000µL), Kimwipes
| Blood Adhesion | Performance Testing | Fill beaker with 10 ml of animal blood. Place medical-grade metal in beaker, place component in beaker, mix by swirling around with hand for 1 minute to simulate blood running along surface of material. Then extract component and medical-grade metal and place component and medical-grade metal in autoclave. Remove from autoclave and take photo of component and medical-grade metal for later reference. | Mix 5 mL PBS and 5 mL of red food dye in large 2000mL beaker. Place medical-grade metal in beaker, place component in beaker, mix by swirling around with hand for 1 minute to simulate blood running along surface of material. Then extract component and medical-grade metal and place component and medical-grade metal in autoclave. Remove from autoclave and take photo of component and medical-grade metal for later reference. | Remove from autoclave and take photo of component and medical-grade metal for later reference. Observe if any red dye is visible, qualitatively compare to amount visible on metal | 2024-11-08 | Equipment and instruments needed Red food dye, Phosphate Buffered Saline (PBS), Motorized pipette filler, Transfusion set components Reusable glassware/plasticware needed Beaker Disposable materials needed Serological pipette,Vacuum pipette
| Biocompatibility | Self-evident by BOM | Compile a comprehensive Bill of Materials (BOM) for the blood-giving sets. Review data sheets for each material used in the design to verify biocompatibility. Cross-check materials against biocompatibility standards. | Compile a comprehensive Bill of Materials (BOM) for the blood-giving sets. Review data sheets for each material used in the design to verify biocompatibility. Cross-check materials against biocompatibility standards. | Compile biocompatibility of all materials and summarize findings in testing document | 2024-11-08 | List of materials and corresponding BOM for each material
2.1 | Durability
| Compressive Strength | Strength Testing | Prepare each component of the blood-giving set for testing by measuring its initial dimensions using calipers. Place the component into the compression testing machine. Apply a gradually increasing compressive force until the component breaks or malfunctions. Record the maximum load at failure for each component. Repeat the test for each individual component of the blood-giving set 3 times each. | Prepare each component of the blood-giving set for testing by measuring its initial dimensions using calipers. Place the component into the compression testing machine. Apply a gradually increasing compressive force until the component breaks or malfunctions. Record the maximum load at failure for each component. Repeat the test for each individual component of the blood-giving set. | Record the maximum load at failure for each component. Repeat the test for each individual component of the blood-giving set. | 2024-11-04 | Materials to be tested including: tubing and 3D printed components. Conduct testing in lab with compression testing machine, report findings.
| Drop Strength | Drop testing | Separate each component from the blood giving set, drop each component onto concrete floor from a 1m height. Repeat 3 times per each component. Note any visible abrasions/damage to the material, take a picture of the material at each level of dropping. | Separate each component from the blood giving set, drop each component onto concrete floor from a 1m height. Repeat 3 times per each component. Note any visible abrasions/damage to the material, take a picture of the material after each round. After damaging piece, put back into set, run water through. If set does not leak after all trials, component passes test. | Record observations for each component for each time dropped and whether or not the material leaks after all the tests. | 2024-11-05 | Transfusion test, phone camera
3.1 | Functional Requirements
| Ease of Use | User Survey | Create an assembly and sterilization guide and have 20 doctors and nurses set up the transfusion set and using a Qualtrics survey rate the ease of use of our design on a likert scale (strongly agree, agree, disagree, strongly disagree) to four questions regarding the instructions, assembly, flow modulation, and disassembly. We also have included a comments section for expansion on their rankings. | Create an assembly and sterilization guide and have 10 classmates set up the transfusion set and using a Qualtrics survey rate the ease of use of our design on a likert scale (strongly agree, agree, disagree, strongly disagree) to four questions regarding the instructions, assembly, flow modulation, and disassembly. We also have included a comments section for expansion on their rankings. | Collect user surveys rating the ease of use of our design on Likert scale (strongly agree, agree, disagree, strongly disagree) | 2024-11-11 | User-survey, timer, and set up instructions to perform the test. Ask 10 students to perform the setup and complete the survey.
| Leakage | Leakage Testing | Set up the transfusion set. Puncture a bag filled with PBS and evaluate and points on leakage in the tubing and connections. Have a bin under the set to quantify unwanted leakage. Have a separate bin to collect PBS from needle tip. Repeat 10 times. | Set up the transfusion set. Puncture a bag filled with PBS and evaluate and points of leakage in the tubing and connections. Have a bin under the set to quantify unwanted leakage. Have a separate bin to collect PBS from needle tip. Repeat 5 times. | Collect PBS from bin under the transfusion set. Measure the water in the bin to quantify the amount of water that leaks out during the test. Average the scores from the trials. | 2024-11-04 | Ziploc with PBS, transfusion set, beakers to perform the test. Repeat test 5 times and record amount of water that leaked.
| Flow Modulation | Flow Testing | Set up the transfusion set. Set the flow modulation at fully open, and let water flow through set for 2 minutes. Collect water from needle tip in a beaker. Weigh water in beaker and using math determine flow rate. Repeat this step in very small increments until fully closed. | Set up the transfusion set. Set the flow modulation at fully open, and let water flow through set for 2 minutes. Collect water from needle tip in a beaker. Weigh water in beaker and using math determine flow rate. Repeat these steps for 75%, 50%, and 25% flow modulation until fully closed. | Collect water from beaker. Determine flow rate at each modulation level using math. Create a report to analyze the modulation | 2024-11-04 | Transfusion set, beaker, water, scale, and math formulas. Alter flow modulation levels and record amount of water in end beaker.
| Filter | Filtration Testing | Mix 250mL PBS with 40mg of 80µm glass beads in 500mL beaker. Hold filter and base above second beaker, pour mixture from first beaker through to second beaker, let it drip for 30 seconds. Hold industry-standard blood filter above third beaker, pour contents from second beaker through and let filter drip for 30 seconds into third beaker. Take contents of industry-standard filter, pour onto plastic holder, ensure that only beads are on scale Determine and record weight of beads filtered out by industry-standard filter that were not filtered out by the filter and base from transfusion set Repeat 5x | Mix 250mL PBS with 40mg of 300µm glass beads in 500mL beaker. Hold filter and base above second beaker, pour mixture from first beaker through to second beaker, let it drip for 30 seconds. Hold industry-standard blood filter above third beaker, pour contents from second beaker through and let filter drip for 30 seconds into third beaker. Take contents of industry-standard filter, pour onto plastic holder, ensure that only beads are on scale Determine and record weight of beads filtered out by industry-standard filter that were not filtered out by the filter and base from transfusion set Repeat 3x | Determine the weight filtered out by the second filter to determine if it is within our target | 2024-11-06 | Equipment and instruments needed Filter and base from blood transfusion set, Industry-standard blood filter and drip chamber, 80µm diameter glass beads. PBS bottle,Gram scale Reusable glassware/plasticware needed Beakers (500mL) Disposable materials needed Plastic holder
| Compact Sizing | Self-evident by dimension of set | Determine the length, width, and height of device in final packaging to be stored on hospital shelfs. Confirm our dimensions are similar in size to current disposable blood giving set used | Determine the length, width, and height of device in a ziploc bag to be stored on hospital shelfs. Confirm our dimensions are similar in size to current disposable blood giving set used | Determine the length, width, and height of device when wrapped up for storage. | 2024-11-06 | Transfusion set Ziploc bag Calipers
4.1 | Community Design Requirements
| Community Comfort | User Survey | Create Qualtrics survey with questions about various community comfort points surrounding our product and in partnership the surgical center, have them hand out a 4 question survey to 100 patients about their impressions of our device during implementation in the surgical center. | Create Qualtrics survey with questions about various community comfort points surrounding our product. Hand out a 4 question survey to 10 classmates about their impressions of our device. | Record responses and critically analyze the trends found in them from the Likert scale of 1 (Strongly disagree) to 5 (Strongly agree) | 2024-11-11 | Create survey with questions to gauge community comfort with our design
| Local Sourcing | BOM & User Survey | Create a list of all materials we plan to use in our design and create a BOM to ensure the materials are locally sourced. Cross check our findings with our client and connections at Malawi Polytechnic | Create a list of all materials we plan to use in our design and survey our partners at Malawi Polytechnic on the local availability of them | Compile a list of our materials and create a document ranking their resource availability on a scale from 1-5 with 5 being the easily found in Malawi | 2024-11-08 | List of materials used in our design. Research whether those materials are available locally in malawi.
| Cost Efficient | Self-evident by BOM | Determine the price of each component used in the transfusion set. Sum up the total cost of the set. | Determine the price of each component used in the transfusion set. Sum up the total cost of the set. | Report cost for each component and total costs of set | 2024-11-08 | List of materials used in our design. BOM of those materials with prices

## Schedule
Test | Design Requirements Being Tested | Things to test/report | Start Date | Target Completion Date | Target Reporting Completion
Bacterial Removal | Safe and Sterilizable | % of bacteria present | 2024-11-04 | 2024-11-06 | 2024-11-10
Blood Adhesion |  | Visibility of red food dye | 2024-11-04 | 2024-11-06 | 2024-11-10
Biocompatibility |  | Biocompatibility according to BOM | 2024-11-05 | 2024-11-05 | 2024-11-10
Compressive Strength | Durable | Maximum load at failure | 2024-11-04 | 2024-11-04 | 2024-11-06
Ease of Use | High Performance | User survey feedback | 2024-11-07 | 2024-11-11 | 2024-11-11
Leakage |  | Amount of leakage observed | 2024-11-04 | 2024-11-04 | 2024-11-06
Flow Modulation |  | Flow rate at different modulation levels of flow regulator | 2024-11-04 | 2024-11-04 | 2024-11-06
Filtration |  | Weight of beads filtered by industry-standard that was not filtered by our device | 2024-11-04 | 2024-11-06 | 2024-11-08
Community Comfort | Compatible with Community | User survey feedback | 2024-11-08 | 2024-11-11 | 2024-11-11
Local Sourcing |  | Local availability according to partners at Malawi Polytechnic | 2024-11-05 | 2024-11-08 | 2024-11-08
Cost Efficient |  | Cost of materials according to BOM | 2024-11-05 | 2024-11-08 | 2024-11-08

## Leakage
Trial | Amount of Unwanted Leakage (ml)
1 | 0
2 | 0
3 | < 1 ml
4 | 0
5 | 0

## Flow Modulation
Water Collected in Beaker
Flow Modulation % | Trial 1 | Trial 2 | Trial 3
100 | 473.2 | 443.6 | 467.3
75 | 360 | 390 | 500
50 | 490 | 470 | 500
25 | 225 | 400 | 400
0 | 0 | 0 | 0